Contents

Summary

- Introduction

- Experimental section

  - Solid phase extraction (SPE)

  - Liquid chromatography –mass spectrometry (LC-MS)

  - Headspace GC-MS analysis for volatile compounds

  - GC-MS analysis for unpolar semivolatile pollutants

  - Physical and sum parameters

- Results and discussion

  - Liquid chromatography –mass spectrometry

  - Solid phase extraction

  - Sequential solid phase extraction

  - Analysis of real water samples

    - LC-MS analysis

    - Physical characterisation and sum parameters

    - C18-SPE CG-MS analysis of unpolar semivolatile contaminants

    - Headspace GC-MS analysis

- Conxlusions and outlook

Tables

Figures

Annexes

Acknoledgments

References

# Summary

The objective of the IPPC Directive 96/61/EC is to prevent or reduce industrial emissions into the environment. Article 14 demands the monitoring of releases affecting the environment.

In support of the IPPC Directive, analysis methods for organic micropollutants in water samples were developed, optimized, and validated. The procedures can be subdivided in different analytical lines:

	Polar water-soluble substances like phenols, acids, sulfonates, dyes, drugs, amines, 	surfactants, pesticides are analysed by solid-phase extraction (SPE) followed by liquid 	chromatography-mass spectrometric (LC-MS) detection.

	Volatile compounds like benzene, toluene, chloroform, methylenechlorid, naphthalene, 	etc. are analysed by headspace gas chromatography/mass spectrometry (GC-MS).

	Unpolar semivolatile contaminants such as unpolar pesticides (DDT), polyaromatic 	hydrocarbons (PAHs), phthalates, chlorobenzenes, hydrocarbons, and PCBs are 	analysed by C18-SPE followed by GC/MS.

	Physical and sum parameters like pH, turbidity, DOC, TOC, color, etc.

For the polar to medium-polar compounds, sample clean-up and group separation are performed with a newly developed sequential SPE procedure using automated SPE with C18 and polymeric sorbent materials. The compounds are divided by this procedure in three different polarity classes. More then 80 standard compounds of environmental interest have been chosen for the experiments. The extraction recovery values were for most of the compounds in the range of 50-100 %. Analyses are performed by negative and positive electrospray ionization (ESI) LC-MS using a single quadrupole system. Due to interfering matrix substances, detection of nonylphenol (NP) in real water samples was possible only after the SSPE procedure.

Moreover, normal SPE of water samples was thoroughly studied for the extraction of dyes, pesticides, phenols, aromatic sulfonates, and the endocrine disrupting alkylphenolic compounds nonylphenol (NP), octylphenol (OP), bisphenol A (BPA), and nonylphenoxy-acetic acid (NPE1C). C18 adsorbs most of the compounds at neutral pH, however very polar substances are only adsorbed by polymeric sorbents. The importance of the choice of the right elution solvent was pointed out. Ethylacetate was the best suited organic solvent for the elution of nonpolar compounds like alkylphenols (APs). In contrast, methanol and acetone are well suited for the elution of polar compounds.

The developed analysis methods have been tested by analysing different water samples like effluents from wastewater treatment plants (WWTPs), and the corresponding receiving waters (rivers and lakes) and the good performance has been shown. The Lariana Depur WWTP „Alto Seveso“ in Fino Mornasco (Italy) showed good treatment performance. Industrial pollution was observed in the Italian rivers Seveso, Olano, Lura, Guisa, and Pudiga.

# 1. Introduction

The IPPC (Integrated Pollution Prevention and Control) Directive 96/61/EC [1] is being implemented in all European countries. This involves the definition of „Best Available Techniques (BATs) for several industrial processes. The objective is to prevent or reduce industrial emissions into the environment. Article 14 demands the monitoring of releases affecting the environment. Annex III gives a indicative list of the main polluting substances. A recent updated list of priority substances in the field of water policy provides Decision No. 2455/2001/EC [2], which is Annex X to the Water Framework Directive 2000/60/EC (see Annex 1) [3].

Environmental trace analysis of priority organic pollutants usually is performed by gas chromatography-mass spectrometry (GC-MS) or liquid chromatography-mass spectrometry (LC-MS) detection techniques after extraction of the samples by liquid-liquid- (LLE) or solid-phase extraction (SPE). LC-MS is increasingly used for the analysis of polar water-soluble compounds [4-9].

Of special interest are polar hydrophilic environmental pollutants because they are readily soluble in water and therefore pose several problems for fresh water systems and the supply of clean drinking water. They are difficult to remove during water treatment processes. In contrast, the more nonpolar water components are usually well removed in wastewater treatment plants (WWTPs) due to their adsorption to soil organic matter.

Prior to the instrumental analysis, the same attention has to be paid to the sample preparation, extraction and clean-up procedures. Nowadays, SPE is the preferred extraction procedure in environmental analytical chemistry. However, the analysis of polar hydrophilic compounds is still in its infancy and not yet routine work, little is know about the environmental behaviour of these compounds [10,11].

Of special interest is the environmental fate of the metabolites of alkylphenol ethoxylate (APEO) surfactants because of the potential effects of these compounds as endocrine disrupters [12-15]. Due to its endocrine disrupting potential, nonylphenol (NP) and octylphenol (OP) have been included in the list of priority substances in the field of water policy which is part of the Water Framework Directive of the European Commission (EC) [2,3].

The objective of this work is the establishment of analysis methods for the monitoring of organic priority micropollutants in the European textile industry and the corresponding receiving waters to access the impact of the industries on the environment. The knowledge of the priority pollutants present in the effluents of textile plants or the corresponding waste water treatment plants is the first step in the development and improvement of the water treatment processes.

# 2. Experimental section

## 2.1. Solid-phase extraction (SPE)

SPE for the concentration of water samples was performed automatically using an AutoTraceTM SPE workstation (Tekmar®, Cincinnati, Ohio, USA). Isolute ENV+ (200 mg, 6 ml), Isolute C18-endcapped (500 mg, 6 ml) from International Sorbent Technology (IST, Cambridge, U.K.), Oasis HLB (200 mg, 6 ml, Waters, Milford, MA, USA), and LiChrolut EN (500 mg, 6 ml, Merck, Darmstadt) SPE cartridges were used.

Simple one-step SPE procedures were performed with C18, Oasis HLB, LiChrolut EN, or Isolute ENV+ cartridges for different analytes. The cartridges were treated as reported beneath for the SSPE procedure. Elution was performed at a flow-rate of 1 ml/min (AutoTrace elute air push 5 ml/min) testing different elution solvents and volumes.

Sequential solid-phase extraction (SSPE) for polar compounds

First, the water is passed at its actual neutral pH through the C18 cartridges, the percolated water is collected with the help of teflon tubes and then this water after addition of 10 mM ion-pairing reagent/acetic acid further extracted by ion-pair SPE using Oasis HLB columns (Figure 1). The ion-pairing reagents used were triethyl- and tributylamine (TEA, TBA). Acidification with 500 µl formic acid (pH 3) resulted in higher recoveries.

Both the C18 and Oasis HLB sorbents are activated and conditioned first with 5 ml methanol and 5 ml water at a flow-rate of 5 ml/min. The sorbents are not allowed to dry, and subsequently the water samples (between 200-400 ml) passed through the cartridges at a flow-rate of 5 ml/min. After this loading step the columns are rinsed with 2 ml water (flow 3 ml/min) for the removal of interfering substances, and then the cartridges completely dried during 30 min using nitrogen at 0.6 bar.

Elution of the C18 cartridges is first performed with 2 x 3 ml hexane/dichloromethane (3:1) giving the first nonpolar fraction. The vials are exchanged and the cartridges subsequently eluted with 2 x 3 ml methanol/acetone/ethylacetate (2:2:1) containing 0.1% formic acid yielding fraction two. The third polar fraction of the water samples is obtained from the Oasis HLB cartridges by elution with 2 x 3 ml methanol/acetone (Figure 1).

The evaporation of the samples with nitrogen is performed at a temperature of 40°C in a water bath using a TurboVap® II Concentration Workstation (Zymark, Hopkinton, MA, USA). The first fraction of the SSPE procedure has to be evaporated to dryness because hexane/dichloromethane is not compatible with LC. The analytes are redissolved in ethylacetate.

C18-SPE procedure for unpolar semivolatile pollutants

Sample pretreatment: The water is adjusted to pH 2 with 6N hydrochloric acid. To remove the particles, the water is decanted into another 1l bottle, followed by the addition of 10 ml 2-propanol and the internal standard (200 µl of Supelco mix 4-8902, 25 µg/l).

Conditioning of the SPE cartridges: The C18 cartridges (IST Isolute C18, EC, 500 mg/6.0 ml) are conditioned with 10 ml 2-propanol and 10 ml Milli-Q water.

Sample extraction: The system lines are rinsed with 50 ml sample, then 950 ml sample are loaded on the cartridges (flow 15 ml/min). After the loading, the cartridges are washed with 10 ml water, and dryed for 15 min with nitrogen.

Elution: The cartridges are eluted 2 times with 2 ml acetone/ethylacetate (1:1, v/v).

Evaporation: With nitrogen to 0.5 ml (Turbovap).

## 2.2. Liquid chromatography-mass spectrometry (LC-MS)

All experiments were done on an Agilent (Waldbronn, Germany) 1100 Series LC-MSD (bench-top single quadrupole system) that was comprised of a binary pump, vacuum degasser, autosampler, thermostated column compartment with column switching valve, diode-array detector and LC-MSD equipped with an orthogonal interface and a standard atmospheric-pressure ionization (API) source. Electrospray ionization (ESI) in the negative and positive mode was used for detection. LC separations were performed with Superspher 100 RP-18 (250 or 125 mm x 2 mm, 4 µm particles) columns of Merck (Darmstadt, Germany). Eluants for most of the separations were (A) water containing 0.1 % acetic acid (pH 3.5) and (B) acetonitrile. For the separation of APEOs, aromatic amines, and aromatic sulfonates special chromatographic conditions have to be used (see below). For the 125 mm column a linear gradient from 20 to 90 % acetonitrile in 30 min and back to 20 % in 5 min at a flow rate of 0.3 ml/min was used, for the 250 mm column the time was 40 min and the flow 0.4 ml/min. The injection volume was 1 or 2 µl, and was performed automatically by the autosampler. The optimized standard MS parameters were: nitrogen drying gas flow 10 L/min; drying gas temperature 350°C; nebulizer pressure 40 psi; capillary potential 3500 V; fragmentation potential 90 or 100 V. The gain was set to 10. Complete system control and data evaluation was done on the Agilent ChemStation for LC-MS. External calibration was used for quantifications.

Aromatic sulfonates were separated with ion-pair chromatography using volatile tributyl-amine (TBA) as ion-pairing reagent. Eluants were (A) water/acetonitrile (80:20) and (B) acetonitrile/water (80:20) both with 5 mM TBA and acetic acid [16]. The gradient used was: 100 % A for 3 min, then to 70 % B in 20 min, to 20 % B in 5 min, back to 100 % A in 5 min, which was held for 3 min for equilibration. A flow of 0.3 ml/min was used also for the 250 mm Superspher RP-18 column in this case.

APEOs were detected as their NH4+ adduct ions (Table 1) [17,18], 20 mM ammonium acetate was added to both the water (A) and the acetonitrile (B) phase. Linear gradient: 20 to 90 % B in 25 min, 90 % B was held for 5 min, back to 20 % B in 5 min, which was held 5 min for equilibration. Column: Superspher RP-18 (250 mm x 2 mm), flow 0.4 ml/min. MS fragmentation potential 60 V.

LC separation of aromatic amines was performed with a Synergi Polar-RP column (Phenomenex, Torrance, CA, USA, 150 x 2 mm, 4 µ). The water phase (A) was buffered with 10 mM ammonium acetate (pH 6.8). Linear gradient: 10 to 80 % B (acetonitrile) in 25 min, 80 % B was held for 5 min, back to 10 % B in 5 min, which was held for 5 min for equilibration; flow 0.3 ml/min.

## 2.3. Headspace GC-MS analysis for volatile compounds

Headspace GC-MS analysis is performed with a HP 6890 GC and a 5973 GC-MSD (Agilent), using a Gerstel (Mülheim an der Ruhr, Germany) CIS 4Plus cooled injection system and a Gerstel multipurpose MPS2 autosampler. A capillary GC column DB-VRX (J&W, 30 m, 0.32 mm i.d., film thickness 1.8 m) is used. The experimental conditions are shown in the following table:

| GC | MS | Gerstel Autosampler | Gerstel injection system |
| --- | --- | --- | --- |
| Temp. start 40ºC hold for 10 min | Mass range 35-500 | Sample volume 1.5 ml | Equilibration time 0.2 min |
| Rate 8ºC/min | Solvent delay 1.0 min | Incubation temp. 60ºC | Initial temp. - 40ºC |
| Final temp. 200ºC | Source temp. 230ºC | Incubation time 20 min | Initial time 0.15 min |
| hold for 5 min | Quad temp. 150ºC | Syringe temp. 70ºC | Rate 12ºC/min |
| Transfer line |  |  | Final temp. 200ºC |
| temp. 230ºC |  |  | Final time 3 min |

Table 6 gives the standard list of 60 volatile priority compounds which are routinely measured by headspace GC-MS. In addition, non-target analysis of unknown compounds is possible by the use of GC-MS spectral libraries.

Sample preparation is very simple for this analysis method. To 15 ml water sample NaCl-salt and the internal standard are added, the vial is put into the autosampler and the sample automatically injected from the headspace into a GC with cooled injection system. The analysis time is 25 min. Quantification is performed by

isotope dilution with deuterated or C13 internal standards. The detection limits are between 0.1 - 0.5 µg/l.

## 2.4. GC-MS analysis for unpolar semivolatile pollutants

GC-MS is performed with a 6890N GC, 7683 autosampler, and a 5973 MSD mass spectrometer (Agilent) using a Gerstel PTV injector. Injection program: 60°C, initial time 0.02 min, rate 12°C/sec, final temp. 350°C, final time 4 min, cryo on (Peltier), equilibration time 0.5 min, splitless time 1 min. The column is a HT-8, 8 % phenyl-polysiloxane-carborane (SGE, Austin, Texas, USA, 25 m, 0.22 mm i.d., 0,25 µm film), carrier gas: helium 12.3 psi, injection volume 1 µl, GC temperature program: 60°C (1 min), rate 10°/min, final temp. 350°C (5 min), GC runtime 30 min, interface 300°C, MS parameters: source 230°C, 70 eV EI ionization, full scan mode 35-500, 3 min solvent delay time.

## 2.5. Physical and sum parameters

The pH, conductivity, turbidity, dissolved oxygen, and the temperature are measured with a multi-parameter water probe U-22 (Horiba, Sulzbach, Germany).

The Dissolved Organic Carbon (DOC) and Total Organic Carbon (TOC) are measured with a TOC-5000 A Total Organic Carbon Analyzer (Shimadzu, Milano, Italy) using the European-DIN norm 1484 [19]. These parameters can be used as sum-monitoring parameters of organic water pollution.

The spectral absorption coefficient SAK (254 nm) and the color at 436 nm are measured with a Cadas 200 UV/VIS- spectrophotometer (Dr. Lange, Duesseldorf, Germany).

# 3. Results and discussion

## 3.1. Liquid chromatography-mass spectrometry

Analyses are performed by reversed-phase liquid chromatography (RP-LC) followed by electrospray ionization (ESI) mass spectrometry (MS) detection using an atmospheric-pressure ionization (API) source in the negative or positive ionization mode.

HPLC is performed with 2 mm i.d. columns (“micro-LC”) with 4 µm particles. The advantage of micro-LC are reduced reagent, solvent, and sample consumption, reduced cost, improved analyses speed, sensitivity and separation efficiency. The flow rate was between 0.25-0.4 ml/min, the injection volume 1-2 µl.

Sensitivity of LC-MS-analysis is dependent on the ionization efficiency which is typically a compound-dependent parameter influenced by the basicity or acidity of any single compound. Table 1 gives the absolute instrument sensitivity of the bench-top quadrupole LC-MS for all single compounds, at an injection volume of 1 µl and a signal-to-noise ratio of approx. 3 (S/N=3). Some phenols like phenol (pKa 9.9), 2,4-dichlorophenol (pKa 7.6), 2,4-dimethylphenol (pKa 10.5), and 2-chlorophenol (pKa 8.1) have due to their relatively low acidity a small ionization efficiency by ESI. Some dyes like Malachite Green or Indigo Blue were not detected by ESI.

The number of solvents and additives that can be applied for LC-MS is limited. Usually, for LC-MS water, acetonitrile and methanol are used. The buffers are ammonium acetate or formate which are adjusted to the preferred pH with ammonium hydroxide or acetic or formic acid [20]. In this work acetonitrile was used as organic phase. With RP-C18 columns the pH of the water phase should be acidic for the protonation of the silanol groups. Usually, 0.1 % acetic acid was used (pH 3.5).

Very polar compounds like aromatic sulfonates have a low retention by normal LC; they are separated by ion-pair chromatography using volatile tributylamine (TBA) as ion-pairing reagent [16].

APEOs form strong complexes with small cations; they are detected as their NH4+ adduct ions (Table 1); 20 mM ammonium acetate was added to the water and acetonitrile phase [17,18]. Quantitative results for individual APEOs should be interpreted with caution because of variations in the oligomer distribution, and different response factors of individual oligomers depending on the number of ethylene oxide groups [21]. NPE1O and OPE1O have a lower sensitivity than NPE2O and OPE2O [22].

LC separation of aromatic amines cannot be performed with conventional C18-HPLC columns. Severe retention problems and peak-tailing of these basic compounds were observed. Special inert or polymeric columns have to been used. A Synergi Polar-RP column showed good separation results (sharp peaks) for these compounds. The water phase is buffered with 10 mM ammonium acetate. The advantage of buffering the mobile phase is maintaining a constant pH during the analysis, which for ionisable compounds is important to obtain reproducible and robust analyses [23,24].

The atmospheric pressure chemical ionization (APCI) interface was checked for the detection of phenols [25-27] and aromatic amines. Only a slight sensitivity difference for some compounds was observed.

The method limit of detection (LOD) or quantification (LOQ) of environmental analytical methods has to be determined by applying the whole analytical extraction and detection procedure to blank samples [28]. The LOD was determined by blanks from the solvents, cartridges and laboratory material used and the baseline-noise of the detector. With the automatic AutoTraceTM SPE workstation acceptable blank-value dependent LODs < 10 ng/l were achieved for the combined SSPE-LC-MS procedure. Elevated blank-values only were observed for the omnipresent linear alkyl benzenesulfonates, LAS (e.g. dodecyl-benzenesulfonate, m/z 325, LOD ~ 0.1 mg/l). The LODs were determined by the formula: LOD = blank-value + 3  RSD of the blank-value.

## 3.2. Solid-phase extraction

SPE was investigated thoroughly for the isolation of phenols, aromatic sulfonates, dyes, and  APs. For the recovery experiments different water types (tap, Milli-Q, and lake water) were used, different SPE cartridges were tested, various organic solvents examined for the elution of the compounds, the elution volume optimized, and the evaporation of the analytes investigated. Moreover, the recoveries were evaluated at different spike concentrations.

The results for the different compound classes are summarised in the following:

- The best sorbents for tracking polar compounds are Oasis HLB and LiChrolut EN.

	Elution of polar compounds like sulfonates or dyes (Table 2) is best performed with polar 	solvents (acetone, methanol), not with ethylacetate.

	Ion-pair SPE is a well suited method for the extraction of aromatic sulfonates. TEA or 	TBA were used as ion-pairing reagent. Higher recoveries were obtained by further 	acidification with formic acid (pH 3).

	Due to the nonpolar character of NP and OP reversed-phase materials (C18) as well as 	polymeric styrol-divinylbenzene (SDB) are well suited sorbents for their extraction (Table 3 + Table 4).

	Adsorption of APs with a lipophilic side-chain is easy, the bigger problem is the 	desorption of the compounds from the cartridge.

	The choice of the best suited organic solvent for elution of the compounds is crucial. It 	could be shown that ethylacetate has a higher eluting power for the lipophilic APs than 	the usually used polar solvents methanol or acetone (Table 3).

	Also mixtures of acetone or methanol with ethylacetate, hexane, or dichloromethane can 	be used. However, the disadvantage of hexane and dichloromethane is that these 	nonpolar solvents are not compatible with LC. Nonpolar solvents are strongly 	recommended for the elution of APs.

The fact that the choice of the right elution solvent is very important for successful SPE methods was also shown for dyes (colorants). Quantitative recoveries were obtained with C18 and Oasis HLB using methanol/acetone elution. With methanol/acetone/ethylacetate the recoveries were significant worse. C18 didn’t adsorb the two most polar compounds, acid red 1 and acid orange 6 (Table 2).

The SPE results for the isolation of APs have been submitted for publication in the Journal of Chrom. A [29].

## 3.3. Sequential solid-phase extraction

Precise analytical results are dependent on a successful combination of sample preparation, chromatographic separation, and analytical detection. Interfering matrix compounds make a thorough sample preparation as well as sophisticated chromatographic separation often essential for reliable quantitative results, also if selective and specific mass spectrometric detection is used.

Ion suppression is a well known LC-MS-effect, which is induced by coeluting matrix components that can have a dramatic effect on the intensity of the analyte signal [20,30,31].

Because of this matrix problems, a sequential solid-phase extraction (SSPE) method was developed for the analysis of water samples (Figure 1). The aim of this procedure is to achieve a group separation (fractionation) of the complex mixture of compounds. First, most of the compounds are adsorbed at neutral pH on a C18 (end-capped) cartridge. Only the most polar substances (e.g. sulfonates, smaller carboxylic acids) are not adsorbed on C18; they are extracted in a second adsorption step with a polymeric adsorbent material (Oasis HLB) by ion-pair extraction using 10 mM tributylamine/acetic acid and 500 µl formic acid (pH 3), yielding the 3rd fraction. Higher recoveries were obtained with the addition of formic acid. The use of end-capped C18-phases ensures that polar compounds completely pass the C18 cartridge. Nonpolar compounds (like NP) are eluted in a 1st fraction from the C18 cartridge using hexane/dichloromethane (3:1) elution. The rest of the compounds are eluted from the C18 cartridge in the 2nd fraction using a methanol/acetone/ethylacetate mixture. The recovery values for all single compounds in the 3 different fractions are given in Table 5. The recoveries were for most of the compounds between 50 – 100 %. The first fraction only contains the most nonpolar compounds NP, OP, 2,4,6-trichlorophenol, PCP, and chlorpyrifos. Most of the compounds are eluted in fraction 2, fraction 3 contains the most polar compounds: aromatic sulfonates, nitrophenols, smaller carboxylic acids, and some dyes.  The smallest most polar aromatic amines were not recovered in the 2nd fraction. They have to be extracted by ion-pair extraction using negative ion-pairing reagents.

The advantage of this SSPE procedure is the group separation of the compounds. The different compound classes are subdivided into 3 different fractions which makes their analytical determination much easier. Otherwise coeluting interfering compounds can be a problem during the chromatographic separation and MS-detection. E.g., successful detection of NP in real water samples was only possible after this SSPE procedure which discriminates interfering matrix components (see Figure 2). APs were eluted in a first nonpolar fraction from C18.

## 3.4. Analysis of real water samples

The performance of the developed procedures was checked by the analysis of different real water samples.

The water samples were not filtered, but decanted to remove the particles. Then, the samples were extracted by the developed SSPE procedure (400 ml for fraction 1 and 2, 300 ml for fraction 3). The different fractions were analysed by LC-ESI-MS in the positive and negative scan-mode, and by positive and negative selected ion monitoring (SIM) detection applying the masses of the known priority substances for which standards were available. A retention time comparison was performed with these standards. A target compound was declared to be identified when the retention time and the peak shape matched exactly with the standard. The reported values are minimal concentrations because they were not corrected with the SPE recovery rates.

### 3.4.1. LC-MS analyses

A first sampling campaign to a Textile WWTP in Fino Mornasco close to Como (Italy) was undertaken on the 19.07.2001. Samples were taken before and after (effluent) the ozonation step of the WWTP. Also the water of the small stream in which the plant discharges was examined before and after the WWTP. The substances which have been found are shown in Annex 2. The concentrations of the compounds in the effluent of the plant were nearly exactly the same as in the river after the plant, as the volume flow of the small stream before the plant was very small. Higher pollutant concentrations were found in the small stream before the WWTP than in the effluent of the plant. This shows direct discharges into the river.

A disadvantage of LC-MS is that the scan-mode is relatively unsensitive. Therefore, only very high concentrated substances can be detected in this mode, e.g. LAS surfactants (see Figure 3).

In the positive scan-mode some peaks were detected, but the masses could not be matched to known substances. The chromatograms showed clearly that the ozonization step destroys some compounds, but also that the WWTP contributes to the pollution of the river. Some substances were formed during the ozonation step, too.

Another sampling campaign was carried out on the 23.10.2001 in the area of Milano (Italy). Eight samples (Annex 3) were taken from different rivers, 7 samples in the north, 1 sample in the south of Milano (NM4). Fraction 1 was wrongly eluted with ethylacetate, not with hexane/dichloromethane. Therefore, NP could not be quantified in these samples. Under this conditions, the interfering matrix background was too high for the detection of NP.

The substances identified in these water samples by SSPE-ESI-LC-MS are reported in Annex 4. Some phenols, aromatic sulfonates, one dye (Uniblue A), pharmaceuticals, LAS- and APEO-surfactants, and the NPEO-metabolite NPE1C were identified. The three most polluted samples were NM4 (south of Milano), NM6, and NM8. Especially high surfactant concentrations have been found in these samples.

In the positive scan mode in sample NM8 the masses 287, 230, 315, 302, 343, 272, 304, 371, and 399 were detected (Figure 4), but could not be identified.

A third sampling campaign was carried out on the 12.02.2002 along the river Seveso north of Milano (Annex 5). RS1 is the sample taken most upstream close to the origin of the river, and RS8 the most downstream sample. Annex 6 gives the results for the polar contaminants found in this river samples. The cleanest water sample was the most upstream sample RS1, however LAS-surfactants were found in this sample. A clear pollution increase was observed more downstream indicating industrial and domestic discharges into the river.

### 3.4.2. Physical characterisation and sum parameters

Annex 7 gives the physical and sum parameters measured in the water samples of the river Seveso. The TOC and DOC values of sample RS1 were significant lower than in the more downstream samples. The highest DOC and TOC values were found in the samples RS 4 and 5.

### 3.4.3. C18-SPE GC-MS analysis of unpolar semivolatile contaminants

Annex 8 gives the analysis results for the unpolar semivolatile compounds found in the water samples of the river Seveso. Only hydrocarbons and phthalates were identified. The rivers Olano, Lura, Guisa, and Pudiga were higher polluted (Annex 9). Many PAHs, chlorinated benzenes, and pesticides were found. The most polluted samples were NM 4, 6 and 8. The results for phthalates have to be interpreted with caution because of laboratory blank problems for this compounds.

### 3.4.4. Headspace GC-MS analyses

Annex 10 and Annex 11 give the results for the volatile contaminants found in the water samples.

# 4. Conclusions and Outlook

In support of the IPPC Directive, analysis methods for organic micropollutants in water samples were developed, optimized, and validated.

A multi-component analysis method for polar to medium-polar water pollutants was developed, consisting of sequential solid-phase extraction (SSPE) followed by liquid chromatography-mass spectrometry (LC-MS) detection. The scope of this method is the possibility to analyse as many compounds and substance classes as possible. The importance of the use of advanced clean-up and SPE procedures was pointed out as the detection of nonylphenol in real water samples only was possible after the SSPE clean-up procedure.

Volatile compounds are routinely analysed by headspace GC-MS, and unpolar semivolatile contaminants by C18-SPE followed by GC-MS. Moreover, physical and sum parameters are measured.

The developed analysis methods have been tested by analysing different water samples like effluents from wastewater treatment plants (WWTPs), and the corresponding receiving waters (rivers and lakes), and the good performance has been shown. The Lariana Depur WWTP in Fino Mornasco (Italy) showed good treatment performance. Industrial and domestic pollution was observed in the rivers Seveso, Olano, Lura, Guisa, and Pudiga.

The work plan for the upcoming work-period is the monitoring and the physical-chemical characterisation of

textile (water treatment plant) effluents and the corresponding receiving waters from the 10 selected textile industries in Italy and Belgium (in cooperation with eco-toxcological characterisations performed by Vito, Belgium) to access their impact on the environment.

Table 1. Priority organic compounds measured by ESI-LC-MS. Column Superspher RP-18 (250 mm x 2 mm), flow 0.4 ml/min; aromatic amines: Synergi Polar-RP column (150 x 2 mm), flow 0.3 ml/min; italic ret. times: ion-pair chromatography with TBA/acetic acid, flow 0.3 ml/min; experimental details: see experimental part.

| # Retention Time | # Compounds | m/z (neg.) | m/z (pos.) | Molecular Weight | Sensitivity [pg] |
| --- | --- | --- | --- | --- | --- |
|  | # Alkylphenolic compounds (EDCs) |  |  |  |  |
| 17.1 | Bisphenol A | 227 |  | 228 | 50 |
| 25.6 | NPE1C, Nonylphenoxyacetic acid | 277 |  | 278 | 20 |
| 25.6 | NPE1C (d2) | 279 |  | 280 | 20 |
| 26.8 | OPE1C, Octylphenoxyacetic acid | 263 |  | 264 | 30 |
| 28.3 | Nonylphenol, technical mixture | 219 |  | 220 | 50 |
| 31.6 | 4-n-Nonylphenol (d8) | 227 |  | 228 | 50 |
| 29.4 | Octylphenol | 205 |  | 206 | 50 |
| 23.8 | OPE1O, Octylphenolmonoethoxylate |  | 268 | 250 | 100 |
| 24.5 | OPE2O, Octylphenoldiethoxylate |  | 312 | 294 | 5 |
| ### 29.0 | ### NPE1O, Nonylphenolmonoethoxylate |  | 282 | 264 | 300 |
| 29.0 | NPE1O (d2) |  | 284 | 266 | 300 |
| 28.0 | NPE2O, Nonylphenoldiethoxylate |  | 326 | 308 | 10 |
|  | # Phenols |  |  |  |  |
| 7.8, 10.4 | Phenol | 93 |  | 94 | 1000 |
| 9.9 | 3-Chlorocatechol | 143 |  | 144 | 50 |
| 11.5, 12.3 | Nitrophenol | 138 |  | 139 | 3 |
| 12.2, 15.2 | 2,4-Dinitrophenol | 183 |  | 184 | 10 |
| 13.3, 16.1 | 2-Chlorophenol | 127 |  | 128 | 5000 |
| 15.4 | # 3,4-Dichlorocatechol | 177/179 |  | 180 | 100 |
| 16.8 | 2-Methyl-4,6-dinitrophenol | 197 |  | 198 | 1 |
| 16.9 | 4-Chloro-3-methylphenol | 141 |  | 142 | 10 |
| 18.1 | 3-tert-Butylphenol | 149 |  | 150 | 50 |
| 19.1 | 2,4-Dichlorophenol | 161 |  | 162 | 500 |
| 20.6 | 2,4,6-Trichlorophenol | 195 |  | 196 | 20 |
| 21.6 | 2,4,6-Tribromophenol | 331, 333 |  | 330 | 30 |
| 24.3 | PCP | 265 |  | 264 | 5 |
|  | 2,4-Dimethylphenol | n.d. |  | 122 | - |
|  | # Hydrochinone | n.d. |  | 110 | - |
|  | # 3,4,5-Trichlorocatechol | n.d. |  | 214 | - |
|  | # 2-Bromophenol | n.d. |  | 173 | - |
|  | # Aromatic Sulfonates |  |  |  |  |
| 6.6 | Benzenesulfonate | 157 |  | 157 | 30 |
| 9.2 | 1,5-Naphthalenedisulfonate | 287 |  | 286 | 200 |
| 9.8 | Anthraquinone-1,5-disulfonate | 367 |  | 366 | 200 |
| 10.7 | 1-Amino-2-naphtholsulfonate | 237 |  | 237 | 200 |
| 10.9 | p-Toluenesulfonate | 171 |  | 171 | 30 |
| 12.0 | 1,3,6-Naphthalenetrisulfonate | 367 |  | 365 | 200 |
| 13.1 | 4-Chlorobenzenesulfonate | 191 |  | 191 | 20 |
| 13.8 | 2-Amino-1-naphthalenesulfonate | 222 |  | 222 | 30 |
| 14.3 | 2-Naphthalenesulfonate | 207 |  | 207 | 10 |
| 16.1 | Anthraquinone-2-sulfonate | 287 |  | 287 | 20 |
| 17.2 | Anthraquinone-2,6-disulfonate | 367 |  | 366 | 200 |
| 16.3 | Dodecylbenzenesulfonate (C12-LAS) | 325 |  | 325 | 20 |
|  | # Pharmaceuticals |  |  |  |  |
| 14.6 | Carbamazepine |  | 237 | 238 | 5 |
| 17.5 | Benzafibrate | 360, 274 |  | 361 | 5 |
| 21.3 | Diclofenac | 294, 250 | 296 | 296 | 5 |
| 21.8 | Ibuprofen | 205, 159 |  | 206 | 5 |
| 23.1 | Gemfibrozil | 249 |  | 250 | 5 |
|  | # Dyes |  |  |  |  |
| 15.8 | Acid Red 1 | 464 |  | 466 | 40 |
| 16.1 | Sulforhodamine G | 529 | 531 | 531 | 35 |
| 10.1, 19.3 | Acid Orange 8 | 341 |  | 343 | 20 |
| 11.6, 21.8 | UniBlue A | 483 |  | 485 | 30 |
| 13.0 | Acid Orange 6 | 293 |  | 295 | 20 |
| 16.3 | Disperse Blue 3 |  | 297 | 298 | 400 |
| 16.9 | 1,4-Diaminoanthrachinone | 237 | 239 | 240 | 20(+) 400(-) |
| 17.8 | Sulforhodamine B | 557 | 559 | 559 | 35 |
| 21.8 | 7-Diethylamino-4-methyl-coumarin |  | 232, 254 | 256 | 2 |
| 23.9 | Disperse Red 1 |  | 315 | 318 | 10 |
| 24.2 | Disperse Blue 14 |  | 267 | 268 | 2 |
|  | Malachite Green | n.d. | n.d. | 348 | - |
|  | Acid Orange 74 | n.d. | n.d. | 423 | - |
|  | Cibacron Brilliant | n.d. | n.d. | 910 | - |
|  | Reactive Black 5 | n.d. | n.d. | 904 | - |
|  | Indigo Blue (syn. Vat Blue) | n.d. | n.d. | 264 | - |
|  | Indigo Carmine | n.d. | n.d. | 422 | - |
|  | # Aromatic Amines |  |  |  |  |
| 2.8 | 1,4-Diaminobenzene |  | 109 | 108 | 50 |
| 3.9 | 2-Aminophenol |  | 110 | 109 | 50 |
| 5.2 | 2,4-Diaminoanisole |  | 139 | 138 | 20 |
| 5.4 | 2,4-Diaminotoluene |  | 123 | 122 | 50 |
| 8.1 | Aniline |  | 94 | 93 | 2000 |
| 12.0 | 2-Amino-4-chlorophenol |  | 144 | 143 | 500 |
| 14.4 | 4-Chloroaniline |  | 128 | 127 | 15000 |
| 14.5 | Benzidine |  | 185 | 184 | 5 |
| 16.9 | 2-Aminonaphthalene |  | 144 | 143 | 100 |
| 16.9 | 3,3’-Dimethylbenzidine |  | 213 | 212 | 2 |
| 17.0 | 3,3'-Dimethoxybenzidine |  | 245 | 244 | 3 |
| 17.4 | N,N-Dimethylaniline |  | 122 | 121 | 500 |
| 17.7 | 3,4-Dichloroaniline |  | 162, 164 | 161 | 15000 |
| 19.2 | 4-Aminobiphenyl |  | 170 | 169 | 100 |
| 20.5 | 3,3’-Dichlorobenzidine |  | 254 | 253 | 100 |
| 20.1 | 4-Aminoazobenzene |  | 198 | 197 | 5 |
|  | # Pesticides |  |  |  |  |
| 8.9 | Atrazine-desethyl |  | 188 | 187 | 5 |
| 13.1 | Bentazone | 239 |  | 240 | 10 |
| 13.3 | Simazine |  | 202 | 206 | 5 |
| 15.0 | 2,4-D | 221, 161 |  | 220/2 | 10 |
| 15.6 | MCPA | 199, 141 |  | 200 | 20 |
| 16.2 | Carbaryl |  | 145 | 202 | 5 |
| 16.3 | Atrazine |  | 216 | 215 | 3 |
| 17.1 | Isoproturon | 205 | 207 | 208 | 2 |
| 17.2 | 2,4,5-T | 255, 195 |  | 254 | 10 |
| 17.2 | Dichlorprop | 233, 161 |  | 234/6 | 20 |
| 17.3 | Diuron | 231, 233 | 233 | 234 | 5 |
| 17.6 | Mecoprop | 213, 141 |  | 214 | 20 |
| 19.2 | Terbutylazine |  | 230 | 229 | 3 |
| 19.4 | Propanil | 216 | 218 | 218 | 5 |
| 20.3 | Linuron |  | 249 | 250 | 5 |
| 22.2 | Alachlor |  | 238, 162 | 270 | 5 |
| 23.3 | Chlorfenvinphos |  | 359 | 358 | 5 |
| 26.3 | Trifluralin |  | 336 | 338 | 50 |
| 28.4 | Chlorpyrifos |  | 352 | 350 | 50 |
| ### NPEO (OPEO) | ### Alkylphenolethoxylate Surfactants |  | (+NH4) |  |  |
| 27.8 (25.2) | NPE3O (OPE3O) |  | 370 (356) | 352 (338) | 500 (300) |
| 27.7 (25.1) | NPE4O (OPE4O) |  | 414 (400) | 396 (382) | 200 (100) |
| 27.6 (25.0) | NPE5O (OPE5O) |  | 458 (444) | 440 (426) | 100 (50) |
| 27.5 (24.9) | NPE6O (OPE6O) |  | 502 (488) | 484 (470) | 100 (50) |
| 27.4 (24.8) | NPE7O (OPE7O) |  | 546 (532) | 528 (514) | 100 (50) |
| 27.3 (24.7) | NPE8O (OPE8O) |  | 590 (576) | 572 (558) | 100 (50) |
| # 27.3 (24.6) | # NPE9O (OPE9O) |  | 634 (620) | 616 (602) | 200 (100) |
| ### 27.2 (24.5) | ### NPE10O (OPE10O) |  | 678 (664) | 660 (646) | 300 (200) |
| 27.1 (24.4) | NPE11O (OPE11O) |  | 722 (708) | 704 (690) | 400 (300) |
| 27.0 (24.3) | NPE12O (OPE12O) |  | 766 (752) | 748 (734) | 500 (500) |
| 26.9 (24.2) | NPE13O (OPE13O) |  | 810 (796) | 792 (778) | 800 (500) |
| 26.9 (24.1) | NPE14O (OPE14O) |  | 854 (840) | 836 (822) | 800 (700) |
| 26.8 (23.9) | NPE15O (OPE15O) |  | 898 (884) | 880 (866) | 800 (800) |

Table 2. Recoveries of dyes from Milli-Q water. Neutral pH, Elution with 2 x 3 ml Methanol/Acetone. (n = 4). Concentration 2-5 µg/l.

Positive Mode

| # Analytes | # C18 | # ENV+ | # Oasis |
| --- | --- | --- | --- |
| Sulforhodamine G | 110 (4) | 23 (6) | 110 (5) |
| Sulforhodamine B | 105 (9) | 40 (0) | 115 (7) |
| 1,4-Diaminoanthrachinone | 105 (2) | 66 (9) | 109 (11) |
| 7-Diethylamino-4-methyl-coumarin | 104 (1) | 98 (13) | 106 (9) |
| Disperse Red 1 | 102 (1) | 32 (6) | 103 (9) |

Negative Mode

| # Analytes | # C18 | # ENV+ | # Oasis |
| --- | --- | --- | --- |
| Acid Red 1 | n.d. | n.d. | 105 (10) |
| Acid Orange 6 | 3 (2) | 15 (2) | 107 (7) |
| Sulforhodamine G | 114 (5) | 23 (1) | 110 (9) |
| Sulforhodamine B | 109 (9) | 33 (2) | 114 (8) |
| Acid Orange 8 | 108 (5) | 32 (3) | 110 (9) |
| UniBlue A | 102 (7) | 24 (3) | 107 (14) |

Table 3. Recoveries of NP, OP, BPA, and NPE1C from 200 ml spiked lake and Milli-Q water with Isolute ENV+. Elution with 2 x 3 ml solvent. (n = 4). Spike concentration 5 µg/l. In brackets: relative standard deviations (RSD).  EA = ethylacetate.

| Elution solvent | EA | MeOH | Hexan/ DCM | EA | EA/-Acetone | Acetone |
| --- | --- | --- | --- | --- | --- | --- |
| Water type | Lake water |  |  | Milli-Q water |  |  |
| BPA | 97 (3) | 29 (3) | - | 100 (7) | 105 (6) | 72 (2) |
| NP | 99 (4) | 64 (17) | 60 (14) | 95 (8) | 99 (8) | 65 (3) |
| OP | 93 (3) | 31 (6) | 49 (7) | 85 (12) | 98 (7) | 62 (3) |
| NPE1C | 6 (2) | 6 (2) | - |  | 41 (6) |  |

Table 4. Recoveries of NP, OP, BPA, and NPE1C from 200 ml spiked lake water. Comparison of Isolute ENV+ with C18, elution with 2 x 3 ml ethylacetate (n = 4). Spike concentration 5 µg/l. In brackets: relative standard deviations (RSD).

| # Analytes | # ENV+ (200 mg) | # C18 (500 mg) |
| --- | --- | --- |
| BPA | 97 (3) | 98 (5) |
| NP | 99 (4) | 90 (5) |
| OP | 93 (3) | 84 (7) |
| NPE1C | 6 (2) | 86 (10) |

Table 5. Recoveries of the SSPE procedure. Normal fond: negative ESI; Underline: positive ESI; Italic: ion-pair chromatography, negative ESI. In brackets: relative standard deviations (RSD), n =  4. [-] means “not detected”.

| # Compound | 1. Fraction | 2. Fraction | 3. Fraction |
| --- | --- | --- | --- |
| BPA | - | 66 (9) | # - |
| NPE1C | - | 62 (10) | - |
| NPE1C d2 | - | 55 (8) | - |
| OPE1C | - | 58 (7) | - |
| Nonylphenol, techn. mix. | 71 (15) | 13 (3) | 4 (1) |
| 4-n-Nonylphenol d8 | 64 (11) | - | - |
| Octylphenol | 57 (12) | 19 (4) | - |
| OPE1O | 9 (5) | 54 (6) | - |
| OPE2O | - | 62 (4) | - |
| NPE1O | 12 (7) | 45 (8) | - |
| NPE1O d2 | 10 (6) | 43 (7) | - |
| NPE2O | - | 48 (6) | - |
| Phenol | - | - | # 76 (9) |
| 3-Chlorocatechol | - | 21 (7) | 24 (4) |
| Nitrophenol | - | 8 (1) 8 (2) | 62 (7) 65 (8) |
| 2,4-Dinitrophenol | - | 1 (0) | 68 (8) 71 (8) |
| 2-Chlorophenol | - | - | - |
| 3,4-Dichlorocatechol | - | 19 (7) | 23 (8) |
| 2-Methyl-4,6-dinitrophenol | - | 7 (3) | 66 (7) |
| 4-Chloro-3-methylphenol | 22 (6) | 45 (7) | 4 (1) |
| 3-tert-Butylphenol | - | 51 (8) | - |
| 2,4-Dichlorophenol | - | - | - |
| 2,4,6-Trichlorophenol | 56 (17) | 14 (3) | 11 (5) |
| PCP | 46 (8) | 33 (4) | - |
| Benzenesulfonate | - | - | 70 (8) |
| 1,5-Naphthalenedisulfonate | - | - | 73 (11) |
| Anthraquinone-1,5-disulfonate | - | - | 60 (6) |
| 1-Amino-2-naphtholsulfonate | - | 72 (15) | - |
| p-Toluenesulfonate | - | - | 74 (8) |
| 1,3,6-Naphthalenetrisulfonate | - | - | 70 (12) |
| 4-Chlorobenzenesulfonate | - | - | 73 (9) |
| 2-Amino-1-naphthalenesulfonate | - | - | 75 (8) |
| 2-Naphthalenesulfonate | - | - | 76 (8) |
| Anthraquinone-2-sulfonate | - | 2 (1) | # 72 (7) |
| Anthraquinone-2,6-disulfonate | - | 52 (16) | - |
| Dodecylbenzenesulfonate | - | 67 (9) | - |
| Carbamazepine | - | 70 (12) | - |
| Benzafibrate | - | 71 (15) | # 13 (9) |
| Diclofenac | - | 72 (13) | 8 (6) |
| Ibuprofen | - | 68 (14) | - |
| Gemfibrozil | - | 69 (9) | - |
| Acid Orange 6 | - | - | 53 (4) |
| Acid Red 1 | - | - | # 79 (10) |
| Sulforhodamine G | - | 70 (12) | # - |
| Sulforhodamine B | - | 69 (12) | - |
| Acid Orange 8 | - | 63 (11) 80 (15) | 3 (1) |
| UniBlue A | - | 61 (9) 76 (12) | - |
| 1,4-Diaminoanthrachinone | - | 73 (9) | - |
| 7-Diethylamino-4-methyl-coumarin | 2 (1) | 67 (11) | - |
| Disperse Red 1 | - | 69 (14) | - |
| Disperse Blue 14 | - | 58 (12) | - |
| 1,4-Diaminobenzene | - | - | # - |
| 2-Aminophenol | - | - | - |
| 2,4-Diaminoanisole | - | - | - |
| 2,4-Diaminotoluene | - | - | - |
| Aniline | - | - | - |
| 2-Amino-4-chlorophenol | - | - | - |
| 4-Chloroaniline | - | - | - |
| Benzidine | - | 53 (5) | - |
| 2-Aminonaphthalene | - | 99 (9) | - |
| 3,3’-Dimethylbenzidine | - | 68 (7) | # - |
| N,N-Dimethylaniline | - | 27 (4) | # - |
| 3,4-Dichloroaniline | - | - | # - |
| 4-Aminobiphenyl | 4 (2) | 62 (5) | # - |
| 3,3’-Dichlorobenzidine | - | 75 (3) | # - |
| 4-Aminoazobenzene | 3 (1) | 79 (6) | # - |
| Atrazine-desethyl | - | 71 (10) | # - |
| Bentazone | - | 1 (1) | 75 (9) |
| Simazine | - | 70 (11) | - |
| 2,4-D | - | 2 (1) | 72 (8) |
| MCPA | - | 3 (1) | 70 (11) |
| Carbaryl | - | 73 (11) | - |
| Atrazine | - | 72 (13) | - |
| Terbutylazine | - | 69 (8) | - |
| Isoproturon | 1 (0) | 69 (7) 72 (13) | - |
| 2,4,5-T | - | 10 (5) | 68 (6) |
| Dichlorprop | - | 10 (5) | 73 (7) |
| Diuron | - | 72 (12) | - |
| Mecoprop | - | 24 (5) | 65 (8) |
| Propanil | - | 72 (12) | - |
| Linuron | 6 (3) | 75 (9) | - |
| Alachlor | 13 (8) | 68 (9) | - |
| Chlorfenvinphos | - | 71 (10) | - |
| Trifluralin | - | 75 (6) | - |
| Chlorpyrifos | 85 (14) | - | - |
| NPE3O (OPE3O) | - | 60 (5) 62 (7) | - |
| NPE4O (OPE4O) | - | 60 (6) 60 (6) | - |
| NPE5O (OPE5O) | - | 57 (8) 60 (7) | - |
| NPE6O (OPE6O) | - | 55 (6) 61 (6) | - |
| NPE7O (OPE7O) | - | 57 (9) 59 (5) | - |
| NPE8O (OPE8O) | - | 56 (4) 59 (7) | - |
| # NPE9O (OPE9O) | - | 57 (8) 59 (7) | - |
| ### NPE10O (OPE10O) | - | 59 (6) 59 (8) | - |
| NPE11O (OPE11O) | - | 58 (7) 58 (7) | - |
| NPE12O (OPE12O) | - | 57 (6) 60 (7) | - |
| NPE13O (OPE13O) | - | 59 (6) 62 (5) | - |
| NPE14O (OPE14O) | - | 64 (7) 58 (6) | - |
| NPE15O (OPE15O) | - | 56 (6) 57 (5) | - |

Table 6. Volatile organic compounds: Headspace analysis in water

by GC-MS.

| No. | Compounds | Detection limit (g/l) | No. | Compounds | Detection limit (g/l) |
| --- | --- | --- | --- | --- | --- |
| 1 | Dichlorofluoromethane | n.d. | 31 | o-Xilene | 0.1 |
| 2 | Chloromethane | 0.1 | 32 | 1,2-Dibromoethane | 0.5 |
| 3 | Vinyl chloride | 0.1 | 33 | 1,3-Dichloropropane | 0.1 |
| 4 | Bromomethane | 1 | 34 | 1,1,2-Trichloroethane | 0.2 |
| 5 | Chloroethane | 0.5 | 35 | trans-1,3 -Dichloropropylene | 0.2 |
| 6 | Trichlorofluoromethane | 0.1 | 36 | cis-1,3-Dichloropropylene | 0.1 |
| 7 | Bromodichloromethane | 0.1 | 37 | 1,1,1,2-Tetrachlorethane | 0.1 |
| 8 | cis-1,2-Dichloroethylene | 0.1 | 38 | 1,1,2,2-Tetrachloroethane | 0.2 |
| 9 | trans-1,2-Dichloroethylene | 0.1 | 39 | Ethylbenzene+m-Xilene | 0.1 |
| 10 | Methylene chloride | 0.1 | 40 | Dibromochloromethane | 0.2 |
| 11 | 1,1-Dichloroethylene | 0.1 | 41 | 1,2,3-Trichloropropane | 0.5 |
| 12 | 1,1-Dichloroethane | 0.1 | 42 | Isopropylbenzene | 0.1 |
| 14 | 2,2-Dichloropropane | 0.1 | 43 | n-Propylbenzene | 0.1 |
| 15 | Chloroform | 0.1 | 44 | 2-Chlorotoluene | 0.1 |
| 16 | 1,1-Dichloropropylene | 0.1 | 45 | 4-Chlorotoluene | 0.1 |
| 17 | 1,2-Dichloroethane | 0.5 | 46 | tert-Butylbenzene | 0.1 |
| 18 | Trichloroethylene | 0.5 | 47 | sec-Butylbenzene | 0.1 |
| 19 | Bromochloromethane | 0.5 | 48 | 1,3-Dichlorobenzene | 0.1 |
| 20 | 1,1,1-Trichloroethane | 0.1 | 49 | 1,4-Dichlorobenzene | 0.1 |
| 21 | Carbon tetrachloride | 0.1 | 50 | 1,2-Dichlorobenzene | 0.1 |
| 22 | 1,2-Dichloropropane | 0.1 | 51 | Bromobenzene | 0.1 |
| 23 | Dibromomethane | 0.5 | 52 | 1,3,5-Trimethylbenzene | 0.1 |
| 24 | Benzene | 0.1 | 53 | 1,2,4-Trimethylbenzene | 0.1 |
| 25 | Toluene | 0.1 | 54 | p-Isopropyltoluene | 0.1 |
| 26 | Tetrachloroethylene | 0.1 | 55 | n-Butylbenzene | 0.1 |
| 27 | Bromoform | 0.5 | 56 | 1,2,4-Trichlorobenzene | 0.1 |
| 28 | Chlorobenzene | 0.1 | 57 | Naphthalene | 0.1 |
| 29 | Styrene | 0.1 | 58 | 1,2,3-Trichlorobenzene | 0.1 |
| 30 | p-Xilene | 0.1 | 59 | 1,2-Dibromo-3-chloropropane | 1 |
|  |  |  | 60 | Hexachlorobutadiene | 0.1 |

Internal Standards: Benzene (d6), Chlorobenzene (d5), Ethylbenzene (d10), 1,4-Dichlorobenzene (d4).

Table 7. Unpolar semivolatile organic contaminants: Analysis in water

by C18-SPE-GC/MS.

| No. | Name | m/z | # Time | No. | Name | m/z | # Time |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Phenol (d6) | 99 | 3.46 | 52 | Dimethylphthalate | 163 | 11.32 |
| 2 | 1,4-Dichlorobenzene (d4) | 150 | 4.42 | 53 | 4-Chlorophenylether | 204 | 12.92 |
| 3 | Naphthalene (d8) | 136 | 7.26 | 54 | N-Nitrosodiphenylamine | 169 | 13.28 |
| 4 | Acenaphtene (d10) | 162 | 11.75 | 55 | Diethylphthalate | 149 | 13.11 |
| 5 | BHC (c13) | 187 | 16.17 | 56 | Azobenzene | 77 | 13.43 |
| 6 | Phenanthrene (d10) | 188 | 15.76 | 57 | 4-Bromophenylether | 248 | 14.20 |
| 7 | DDT (c13) | 247 | 19.75 | 58 | Carbazole | 167 | 16.43 |
| 8 | Chrysene (d12) | 240 | 23.00 | 59 | Di-n-butylphthalate | 149 | 17.27 |
| 9 | Perylene (d12) | 264 | 26.78 | 60 | Butylbenzylphthalate | 149 | 21.22 |
| 10 | Phenol | 94 | 3.40 | 61 | Bis(2-ethylhexyl)phthalate | 149 | 22.50 |
| 11 | 2-Chlorophenol | 128 | 3.97 | 62 | Di-n-octylphthalate | 149 | 24.15 |
| 12 | Aniline | 93 | 3.97 | 63 | 1,3-Dichlorobenzene | 146 | 4.35 |
| 13 | 2-Methylphenol | 107 | 4.63 | 64 | 1,4-Dichlorobenzene | 146 | 4.40 |
| 14 | Benzyl alcohol | 79 | 4.63 | 65 | 1,2-Dichlorobenzene | 146 | 4.45 |
| 15 | 4-Methylphenol | 107 | 4.93 | 66 | Hexachlorethane | 117 | 5.05 |
| 16 | Benzoic acid | 105 | 7.03 | 67 | Nitrobenzene | 123 | 6.23 |
| 17 | 2,4-Dimethylphenol | 122 | 6.32 | 68 | Isophorone | 82 | 6.95 |
| 18 | 2,4-Dichlorophenol | 162 | 6.67 | 69 | 1,2,3-Trichlorobenzene | 180 | 7.10 |
| 19 | 4-Chloroaniline | 127 | 7.32 | 70 | Hexachlorobutadiene | 225 | 7.18 |
| 20 | 4-Chloro-3-methylphenol | 107 | 8.28 | 71 | Hexachlorocyclopentadiene | 237 | 9.10 |
| 21 | 2,4,5 or 2,4,6-Trichlorophenol | 196 | 9.60 | 72 | 2,6-Dinitrotoluene | 148 | 11.78 |
| 22 | 2- or 3- or 4-Nitroaniline | 138 | 11.10 | 73 | 2,4-Dinitrotoluene | 165 | 13.25 |
| 23 | 4-Nitrophenol | 139 | 12.23 | 74 | Hexachlorobenzene | 284 | 14.78 |
| 24 | 2/3/4 Nitroaniline + Dibenzofuran | 168 | 12.23 | 75 | alpha-BHC | 181 | 14.65 |
| 25 | 2,4-Dinitrophenol | 184 | 15.77 | 76 | gamma-BHC | 181 | 15.53 |
| 26 | 2-Methyl-4,6-dinitrophenol | 198 | 15.77 | 77 | beta-BHC | 181 | 15.70 |
| 27 | 2- or 3- or 4-Nitroaniline | 138 | 14.23 | 78 | delta-BHC | 181 | 16.37 |
| 28 | Pentachlorophenol | 266 | 15.23 | 79 | Heptachlor | 100 | 16.65 |
| 29 | Benzidine | 184 | 18.55 | 80 | Aldrin | 66 | 17.18 |
| 30 | 3,3'-Dichlorobenzidine | 252 | 22.52 | 81 | Heptachlor (isomer B) | 81 | 18.40 |
| 31 | Naphthalene | 128 | 7.32 | 82 | ## Endosulfan1 ( | 195 | 19.20 |
| 32 | 2-Methylnaphthalene | 141 | 9.05 | 83 | 4,4-DDE | 246 | 19.75 |
| 33 | Acenaphthylene | 152 | 11.42 | 84 | Endrin | 79 | 20.42 |
| 34 | Acenaphthene | 153 | 11.83 | 85 | ## Endosulfan2 ( | 202 | 20.82 |
| 35 | Fluorene | 165 | 13.18 | 86 | Dieldrine | 79 | 19.81 |
| 36 | Phenanthrene | 178 | 15.83 | 87 | 4,4-DDD | 235 | 20.97 |
| 37 | Anthracene | 178 | 15.93 | 88 | 4,4-DDT | 235 | 21.58 |
| 38 | Fluoranthene | 202 | 19.15 | 89 | Endosulfan sulfate | 272 | 21.97 |
| 39 | Pyrene | 202 | 19.73 | 90 | Methoxychlor | 227 | 22.50 |
| 40 | Benzo(a)anthracene | 228 | 22.96 | 91 | Endrin aldehyde | 345 | 22.52 |
| 41 | Chrysene | 228 | 23.07 | 92 | Endrine ketone | 317 | 23.20 |
| 42 | Benzo(b or k)fluoranthene | 252 | 25.86 | 93 | 2-PCB | 188 | 11.70 |
| 43 | Benzo(a)pyrene | 252 | 26.62 | 94 | 2,3-PCB | 222 | 14.45 |
| 44 | Dibenzo(a,h)anthracene | 278 | 28.98 | 95 | 2,4'-PCB | 222 | 17.43 |
| 45 | Benzo(ghi)perylene | 276 | 28.98 | 96 | 2,4',5-PCB | 256 | 16.17 |
| 46 | Indeno(1,2,3-cd)pyrene | 276 | 29.12 | 97 | 2,2'5-PCB | 256 | 16.17 |
| 47 | Bis(2-chloroethyl)ether | 93 | 4.10 | 98 | 2,4,4'-PCB | 256 | 16.17 |
| 48 | Bis(2-chloroisopropyl)ether | 45 | 5.18 | 99 | 2,2',4,4'-PCB | 292 | 17.43 |
| 49 | N-Nitroso-di-n-propylamine | 130 | 6.22 | 100 | 2,2',3',4,6-PCB | 326 | 18.53 |
| 50 | Bis(2-chloroethoxy)methane | 93 | 6.70 | 101 | 2,2',4,4',5,6'-PCB | 360 | 19.86 |
| 51 | 2-Chloronaphthalene | 162 | 10.30 | 102 | 2,2',3',5,5',6-PCB | 360 | 19.86 |
| No. | Name | m/z | Time | No. | Name | m/z | Time |
| 103 | 2,2',3,3',4,5',6,6'-PCB | 430 | 22.33 | 118 | Hydrocarbon C20 | 71 | 16.93 |
| 104 | 2,2',3,3',4,4',6-PCB | 394 | 22.50 | 119 | Hydrocarbon C21 | 71 | 17.87 |
| 105 | 2,2',3,4,4'5',6-PCB | 394 | 22.50 | 120 | Hydrocarbon C22 | 71 | 18.78 |
| 106 | 2,2',3,4',5,5',6-PCB | 394 | 22.50 | 121 | Hydrocarbon C23 | 71 | 19.69 |
| 107 | 2,2',3,4,4',5,6'-PCB | 394 | 22.50 | 122 | Hydrocarbon C24 | 71 | 20.75 |
| 108 | Hydrocarbon C10 | 71 | 3.85 | 123 | Hydrocarbon C25 | 71 | 21.35 |
| 109 | Hydrocarbon C11 | 71 | 5.31 | 124 | Hydrocarbon C26 | 71 | 22.43 |
| 110 | Hydrocarbon C12 | 71 | 6.85 | 125 | Hydrocarbon C27 | 71 | 22.92 |
| 111 | Hydrocarbon C13 | 71 | 8.50 | 126 | Hydrocarbon C28 | 71 | 24.00 |
| 112 | Hydrocarbon C14 | 71 | 9.75 | 127 | Hydrocarbon C29 | 71 | 24.70 |
| 113 | Hydrocarbon C15 | 71 | 11.45 | 128 | Hydrocarbon C30 | 71 | 25.50 |
| 114 | Hydrocarbon C16 | 71 | 12.38 | 129 | Hydrocarbon C31 | 71 | 26.00 |
| 115 | Hydrocarbon C17 | 71 | 13.60 | 130 | Hydrocarbon C32 | 71 | 26.89 |
| 116 | Hydrocarbon C18 | 71 | 14.77 | 131 | Hydrocarbon C33 | 71 | 27.52 |
| 117 | Hydrocarbon C19 | 71 | 15.99 | 132 | Hydrocarbon C34 | 71 | 28.16 |

The chemical structures of the compounds can be found under

http://chemfinder.cambridgesoft.com/

Figure 1. Sequential SPE procedure.

Figure 2. ESI-LC-MS chromatograms of a river SPE extract and NP standards. (A) Simple SPE extraction with Oasis HLB, elution with methanol/acetone. (B) Corresponding NP standard. (C) First fraction of SSPE with C18, elution with hexane/DCM (3:1). (D) Corresponding NP standard. SIM detection of nonylphenolate [M-H]-, m/z 219. Conditions see  experimental part. Chromatograms (A) and (B) were recorded with a faster gradient than (C) and (D), this explains the retention time difference.

Figure 3. ESI-LC-MS negative scan mode chromatogram of a river SPE extract (river Seveso). Detection of Naphthalenemonosulfonate (m/z 207), C11-LAS (m/z 311), C12-LAS (m/z 325), C13-LAS (m/z 339).

Figure 4. ESI-LC-MS positive scan mode chromatogram of a river SPE extract.  Sample (NM8).

Annex 1:	List of priority substances in the field of water policy (*)

|  | CAS number | EU number | Name of priority substance | Identified as priority hazardous substance |
| --- | --- | --- | --- | --- |
| (1) | 15972-60-8 | 240-110-8 | Alachlor |  |
| (2) | 120-12-7 | 204-371-1 | Anthracene | (X)*** |
| (3) | 1912-24-9 | 217-617-8 | Atrazine | (X)*** |
| (4) | 71-43-2 | 200-753-7 | Benzene |  |
| (5) | n.a. | n.a. | Brominated diphenylethers (**) | X**** |
| (6) | 7440-43-9 | 231-152-8 | Cadmium and its compounds | X |
| (7) | 85535-84-8 | 287-476-5 | C10-13-chloroalkanes (**) | X |
| (8) | 470-90-6 | 207-432-0 | Chlorfenvinphos |  |
| (9) | 2921-88-2 | 220-864-4 | Chlorpyrifos | (X)*** |
| (10) | 107-06-2 | 203-458-1 | 1,2-Dichloroethane |  |
| (11) | 75-09-2 | 200-838-9 | Dichloromethane |  |
| (12) | 117-81-7 | 204-211-0 | Di(2-ethylhexyl)phthalate (DEHP) | (X)*** |
| (13) | 330-54-1 | 206-354-4 | Diuron | (X)*** |
| (14) | 115-29-7 | 204-079-4 | Endosulfan | (X)*** |
|  | 959-98-8 | n.a. | (alpha-endosulfan) |  |
| (15) | 206-44-0 | 205-912-4 | Fluoroanthene(*****) |  |
| (16) | 118-74-1 | 204-273-9 | Hexachlorobenzene | X |
| (17) | 87-68-3 | 201-765-5 | Hexachlorobutadiene | X |
| (18) | 608-73-1 | 210-158-9 | Hexachlorocyclohexane | X |
|  | 58-89-9 | 200-401-2 | (gamma-isomer, Lindane) |  |
| (19) | 34123-59-6 | 251-835-4 | Isoproturon | (X)*** |
| (20) | 7439-92-1 | 231-100-4 | Lead and its compounds | (X)*** |
| (21) | 7439-97-6 | 231-106-7 | Mercury and its compounds | X |
| (22) | 91-20-3 | 202-049-5 | Naphthalene | (X)*** |
| (23) | 7440-02-0 | 231-111-4 | Nickel and its compounds |  |
| (24) | 25154-52-3 | 246-672-0 | Nonylphenols | X |
|  | 104-40-5 | 203-199-4 | (4-(para)-nonylphenol) |  |
| (25) | 1806-26-4 | 217-302-5 | Octylphenols | (X)*** |
|  | 140-66-9 | n.a. | (para-tert-octylphenol) |  |
| (26) | 608-93-5 | 210-172-5 | Pentachlorobenzene | X |
| (27) | 87-86-5 | 201-778-6 | Pentachlorophenol | (X)*** |
| (28) | n.a. | n.a. | Polyaromatic hydrocarbons | X |
|  | 50-32-8 | 200-028-5 | (Benzo(a)pyrene), |  |
|  | 205-99-2 | 205-911-9 | (Benzo(b)fluoroanthene), |  |
|  | 191-24-2 | 205-883-8 | (Benzo(g,h,i)perylene), |  |
|  | 207-08-9 | 205-916-6 | (Benzo(k)fluoroanthene), |  |
|  | 193-39-5 | 205-893-2 | (Indeno(1,2,3-cd)pyrene) |  |
| (29) | 122-34-9 | 204-535-2 | Simazine | (X)*** |
| (30) | 688-73-3 | 211-704-4 | Tributyltin compounds | X |
|  | 36643-28-4 | n.a. | (Tributyltin-cation) |  |
| (31) | 12002-48-1 | 234-413-4 | Trichlorobenzenes | (X)*** |
|  | 120-82-1 | 204-428-0 | (1,2,4-Trichlorobenzene) |  |
| (32) | 67-66-3 | 200-663-8 | Trichloromethane (Chloroform) |  |
| (33) | 1582-09-8 | 216-428-8 | Trifluralin | (X)*** |

*	Where groups of substances have been selected, typical individual representatives are listed as indicative parameters (in brackets and without number). The establishment of controls will be targeted to these individual substances, without prejudicing the inclusion of other individual representatives, where appropriate.

**	These groups of substances normally include a considerable number of individual compounds. Presently, appropriate indicative parameters cannot be given.

***	This priority substance is subject to a review for identification as possible “priority hazardous substance”. The Commission will make a proposal to the European Parliament and Council for its final classification not later than 12 months after adoption of this list. The timetable laid down in Article 16 of Directive 2000/60/EC for the Commission's proposals of controls is not affected by this review..

****	Only Pentabromobiphenylether (CAS number 32534-81-9)

*****	Fluoranthene is on the list as an indicator of other, more dangerous Polyaromatic Hydrocarbons

Annex 2. First sampling campaign to the Textile WWTP “Alto Seveso” in Fino Mornasco. 19.07.2001. Compounds found in a small stream (river Seveso) before the WWTP and the effluent of the textile WWTP discharging into this stream. Concentrations in [μg/l].

| Compound | Effluent | ## River |
| --- | --- | --- |
| Linear alkylbenzenesulfonates (LAS) | n.d. | 14 |
| Mononitrophenol | 0.074 | 0.082 |
| 2,4-Dinitrophenol | n.d. | 0.003 |
| Nonylphenol (NP) | 0.072 | 0.174 |
| NPE1C (and other NPEO metabolites) | 1.2 | 1.5 |
| 4-Chlorobenzenesulfonate | 0.062 | 0.035 |
| Naphthalenemonosulfonate | 0.056 | 0.739 |
| Acid Orange 8 | n.d. | 0.2 |
| Uniblue A | n.d. | 0.074 |
| Sulforhodamine B | 0.133 | 0.15 |

Annex 3: Sampling campaign to the aera of Milano. 23.10.2001.

| Sample Name | Place/City | River | ## CPS coordinates Latitude Longitude |  | Description | pH |
| --- | --- | --- | --- | --- | --- | --- |
| NM1 | Legnano | Olano | 45.353 | 8.552 | less particles | 7.95 |
| NM2 | Nerviano | Olano | 45.332 | 8.584 | less particles | 8.07 |
| NM3 | Pogliano Milanese | Olano | 45.325 | 8.594 | foam | 8.06 |
| NM4 | Milano | Olano | 45.259 | 9.087 | many particles, dirty | 8.00 |
| NM5 | Lainate | Lura | 45.349 | 9.021 | less particles | 8.06 |
| NM6 | Garbagnate | Guisa | 45.351 | 9.048 | disperse | 7.87 |
| NM7 | Bollate | Canale C.S.N.O. | 45.328 | 9.059 | less particles | 8.09 |
| NM8 | Bollate | Pudiga | 45.321 | 9.072 | disperse, oily, dirty | 7.70 |

Annex 4. Sampling campaign to the aera of Milano. 23.10.2001. Polar compounds found in the different river samples by SSPE-LC-MS analysis. Concentrations in [μg/l].

| Compound |  | ## NM1 | NM2 | NM3 | ## NM4 | NM5 | NM6 | NM7 | NM8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2,4,6-Trinitrophenol |  | 0.27 | 0.30 | 0.31 | 0.41 | 0.47 | 0.58 | 0.42 | 0.35 |
| 2,4-Dinitrophenol |  | n.d. | n.d. | n.d. | n.d. | n.d. | 1.2 | n.d. | 2.4 |
| Bisphenol A |  | 0.32 | 0.28 | 0.17 | 0.10 | 0.02 | 0.11 | 0.03 | 0.16 |
| Nonylphenol (1) |  |  |  |  |  |  |  |  |  |
| NPE1C |  | 0.9 | 1.3 | 1.4 | 1.6 | 1.7 | 2.9 | 2.0 | 0.14 |
| Uniblue A |  | 1.1 | 0.60 | 0.30 | 0.04 | n.d. | n.d. | n.d. | n.d. |
| 2-Amino-1-NS |  | 0.19 | 0.46 | 0.65 | 0.14 | n.d. | n.d. | 0.12 | n.d. |
| p-Toluene-sulfonate |  | n.d. | n.d. | n.d. | 0.13 | n.d. | 1.03 | 0.27 | 1.06 |
| Naphthalenesulfonate |  | 4.8 | 4.6 | 4.9 | 0.67 | n.d. | 0.48 | 0.54 | 0.42 |
| LAS (Dodecylbenzenesulfonate) | C10-LAS C11-LAS C12-LAS C13-LAS | 12 19 10 5 | 15 23 14 7 | 21 32 18 9 | 80 157 121 81 | 8 14 9 5 | 192 360 281 178 | 31 51 33 22 | 197 397 295 201 |
| Σ OPEOs (OPE4O) |  | 11 | 10 | 1 | 61 | 4 | 30 | 49 | 124 |
| Σ NPEOs |  | 27 | 19 | 20 | 200 | 15 | 397 | 27 | 925 |
| Benzafibrate |  | 0.034 | 0.024 | 0.024 | 0.093 | 0.016 | 0.117 | 0.047 | 0.051 |
| Gemfibrozil |  | 0.061 | 0.057 | 0.061 | 0.326 | n.d. | 0.086 | 0.084 | 0.712 |
| Diclofenac |  | 0.022 | 0.030 | 0.026 | 0.074 | 0.025 | 0.064 | 0.037 | 0.074 |
| Carbamazepine |  | 0.052 | 0.091 | 0.078 | 0.265 | 0.056 | 0.253 | 0.137 | 0.463 |

- Nonylphenol could not be quantified in these samples because the first fraction was wrongly eluted from the C18 cartridge with ethylacetate, not with hexane/dichloromethane. Under these conditions, the interfering matrix background was too high for the detection of NP.

- Identification and quantification of APEOs features several problems and uncertainties due to the complexity of the mixtures and the lack of pure reference compounds [17,18,21,32].

Annex 5:	Sampling campaign to the River Seveso. 12.02.2002.

| Location | ## Sample name | Time | ## CPS coordinates Latitude Longitude |  | Depth [m] |
| --- | --- | --- | --- | --- | --- |
| Montano Lucino | RS1 | 10.10 | 45.476 | 9.028 | 0.05 |
| Fino Mornasco | RS2 | 10.44 | 45.446 | 9.039 | 0.1 |
| Asnago/Cantu' | RS3 | 11.30 | 45.431 | 9.057 | 0.1 |
| Lentate sul Seveso | RS4 | 11.55 | 45.407 | 9.074 | 0.2 |
| Seveso | RS5 | 14.00 | 45.394 | 9.083 | 0.2 |
| Cesano Maderno | RS6 | 14.24 | 45.631 | 9.143 | 0.2 |
| Bovisio Masnago | RS7 | 14.45 | 45.37 | 9.087 | 0.2 |
| Varedo | RS8 | 15.11 | 45.597 | 9.154 | 0.2 |

Annex 6: Sampling campaign to the River Seveso. 12.02.2002. Concentrations for the polar contaminants measured by SSPE-LC/MS, [μg/l].

| Compound |  | ## RS1 | RS2 | RS3 | ## RS4 | RS5 | RS6 | RS7 | RS8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2,4,6-Trinitrophenol |  | - | - | - | - | - | - | - | 0.225 |
| 2,4-Dinitrophenol |  | - | - | - | - | - | - | 0.046 | 0.061 |
| Mononitrophenol | (3. fr.) | 0.026 | 0.186 | 0.094 | 0.053 | 0.065 | 0.093 | 0.054 | 0.051 |
| Bisphenol A |  | - | - | - | - | - | - | 0.037 | 0.039 |
| Nonylphenol |  | 0.097 | 0.205 | 0.154 | 0.429 | 0.454 | 0.319 | 0.314 | 0.302 |
| NPE1C |  | 0.3 | 2.4 | 1.5 | 3.0 | 3.6 | 2.8 | 2.4 | 3.1 |
| 2-Amino-1-NS |  | - | - | 0.046 | 0.100 | 0.095 | 0.112 | 0.079 | 0.098 |
| p-Toluenesulfonate |  | - | - | - | - | - | 0.502 | 0.355 | - |
| Naphthalenesulfonate |  | - | 0.242 | 0.112 | 0.412 | 0.469 | 0.848 | 0.416 | 0.222 |
| Naphthalene-1,5-disulfonate |  | - | 3.7 | 1.2 | 1.6 | 2.0 | 1.2 | 0.8 | 0.6 |
| LAS (Dodecylbenzenesulfonate) | C10-LAS C11-LAS C12-LAS C13-LAS Σ LAS | 21.3 18.0 43.2 23.9 106.4 | 13.6 11.3 13.7 14.9 53.5 | 7.4 5.6 4.1 6.3 23.4 | 15.8 9.1 6.1 6.6 37.6 | 29.4 19.1 14.0 26.3 88.8 | 135.4 95.7 89.9 135.4 456.4 | 144.8 94.6 72.3 102.1 413.8 | 140.0 90.5 68.2 102.1 400.8 |
| NPE2O (single Standard) |  | - | - | - | 3.2 | 3.4 | 1.3 | 0.6 | 0.6 |
| NPE3O |  | - | - | - | 9.3 | 9.7 | 3.5 | 1.4 | 2.2 |
| Σ NPEOs |  | 9.1 | 12.6 | 2.6 | 236.8 | 241.8 | 124.9 | 66.4 | 99.2 |
| OPE3O |  | - | - | - | - | - | - | 3.9 | 18 |
| OPE4O |  | - | - | - | 36.7 | 40.1 | 24.4 | 21.0 | 27.5 |
| Σ OPEOs |  | - | - | - | 36.7 | 40.1 | 24.4 | 26.3 | 57.7 |
| Benzafibrate |  | - | 0.041 | - | 0.056 | 0.076 | 0.044 | 0.046 | 0.045 |
| Diclofenac |  | - | - | - | 0.037 | 0.041 | 0.022 | 0.016 | 0.021 |
| Carbamazepine |  | - | - | - | 0.056 | 0.064 | - | - | - |

Annex 7: 	Sampling campaign to the River Seveso. 12.02.2002. Physical and sum parameters.

| Compound | ## RS1 | RS2 | RS3 | ## RS4 | RS5 | RS6 | RS7 | RS8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| pH | 5.98 | 7.03 | 8.00 | 7.72 | 7.74 | 7.86 | 7.88 | 7.94 |
| Conductivity [S/m] | 0.04 | 0.09 | 0.08 | 0.14 | 0.15 | 0.10 | 0.10 | 0.10 |
| Turbidity [NTU] | 11.3 | 69.3 | 13.5 | 29.8 | 298 | 49.7 | 234 | 48.4 |
| Disolved oxygen [%] | 42.2 | 100.0 | 61.4 | 81.5 | 86.1 | 80.9 | 83.0 | 83.0 |
| Temperature [ºC] | 4.5 | 11.5 | 6.4 | 8.7 | 9.4 | 8.9 | 8.9 | 8.9 |
| TOC [mg/l] | 2.9 | 10.9 | 6.9 | 16.4 | 15.1 | 10.9 | 14.6 | 9.3 |
| DOC [mg/l] | 2.2 | 8.2 | 4.9 | 9.2 | 9.4 | 8.3 | 6.9 | 7.7 |
| SAK [m-1] 254 nm | 1.32 | 3.87 | 2.61 | 5.12 | 4.98 | 5.14 | 5.27 | 4.14 |
| Color [m-1] 436 nm | 0.132 | 0.515 | 0.31 | 0.915 | 0.852 | 1.1 | 1.27 | 0.775 |

Annex 8: Unpolar semivolatile compounds. 12.02.2002. C18-SPE GC-MS results of the sampling campaign to the River Seveso. Concentrations in [μg/l].

| Compound | ## RS1 | RS2 | RS3 | ## RS4 | RS5 | RS6 | RS7 | RS8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hydrocarbon C10 | 0.60 | 0.88 | 2.29 | 1.04 | 0.44 | 0.44 | 1.84 | 1.99 |
| Hydrocarbon C11 | - | - | 0.39 | 0.28 | 0.19 | 0.19 | 0.45 | 0.44 |
| Hydrocarbon C14 | 0.24 | 0.16 | 0.20 | 0.80 | 1.11 | 1.11 | 0.25 | 0.26 |
| Hydrocarbon C16 | - | - | - | - | - | - | 0.24 | 0.23 |
| Hydrocarbon C17 | 1.24 | 0.49 | 1.53 | 1.35 | 0.73 | 0.73 | 1.11 | 1.60 |
| Hydrocarbon C18 | - | - | - | 0.35 | - | - | - | 0.29 |
| Hydrocarbon C20 | - | 0.39 | 0.31 | 0.18 | - | - | - | 0.14 |
| Hydrocarbon C21 | - | 0.24 | 0.20 | 0.19 | 1.11 | 1.11 | 1.03 | 0.50 |
| Hydrocarbon C23 | - | 0.59 | - | 0.58 | 0.58 | 0.58 | 0.55 | - |
| Hydrocarbon C24 | - | 0.59 | 0.64 | 0.59 | 0.58 | 0.58 | 0.59 | 0.58 |
| Hydrocarbon C25 | - | 0.81 | 0.68 | 0.84 | 0.69 | 0.69 | 0.68 | 0.81 |
| Hydrocarbon C26 | 0.75 | 0.88 | 0.85 | 1.16 | 1.31 | 1.31 | 1.36 | 1.29 |
| Hydrocarbon C27 | 0.85 | 0.91 | 0.76 | 2.30 | 1.36 | 1.36 | 1.11 | 1.80 |
| Hydrocarbon C28 | 1.00 | - | - | 0.89 | 0.91 | 0.91 | - | 2.51 |
| Hydrocarbon C29 | 1.05 | 0.84 | - | - | 1.04 | 1.04 | 0.91 | 2.63 |
| Hydrocarbon C30 | 1.04 | - | - | - | 1.25 | 1.25 | 1.04 | 2.54 |
| Hydrocarbon C31 | - | 1.03 | 0.99 | 2.41 | 1.83 | 1.83 | 1.53 | 1.78 |
| Hydrocarbon C32 | - | 1.10 | 1.06 | 2.85 | 2.03 | 2.03 | 0.96 | 1.63 |
| Hydrocarbon C33 | - | - | - | - | 1.25 | 1.25 | 1.05 | 1.11 |
| Hydrocarbon C34 | - | - | - | 1.21 | 2.15 | 2.15 | 1.81 | 1.60 |
| Di-n-butylphthalate | 1.79 | 1.65 | 2.35 | 2.75 | 2.33 | 2.33 | 2.46 | 2.14 |
| Diethylphthalate | 0.64 | 0.51 | 0.59 | 0.71 | 3.11 | 3.11 | 2.43 | 1.54 |
| Dimethylphthalate | - | 0.09 | - | - | 0.10 | 0.10 | 0.14 | 0.13 |
| Butylbenzylphthalate | - | 1.46 | - | - | 139.2 | 139.2 | 69.9 | 142.6 |
| Bis(2-ethylhexyl)phthalate | 2.14 | 3.18 | 3.64 | 11.36 | 14.56 | 14.56 | 16.91 | 8.73 |
| Di-n-octylphthalate | - | 0.90 | - | 1.36 | 1.90 | 1.90 | 1.63 | 1.73 |

Annex 9: Unpolar semivolatile compounds. 23.10.2001. C18-SPE GC-MS results of the sampling campaign to the area of Milano. Concentrations in [μg/l].

| Compound | ## NM1 | NM2 | NM3 | ## NM4 | NM5 | NM6 | NM7 | NM8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4-Chloroaniline | - | - | - | - | - | - | - | 0.30 |
| Acenaphthylene | - | - | - | - | - | 0.36 | - | 0.23 |
| Acenaphthene | - | - | - | 0.04 | - | 0.09 | - | 0.34 |
| Fluorene | - | - | - | - | - | - | - | 0.56 |
| Phenanthrene | 0.15 | 0.09 | 0.03 | 0.12 | 0.05 | 0.35 | 0.03 | 2.07 |
| Anthracene | 0.15 | 0.10 | 0.03 | 0.12 | 0.05 | 0.37 | 0.03 | 2.15 |
| Fluoranthene | - | - | - | - | - | 0.06 | - | 0.56 |
| Pyrene | - | - | - | - | - | 0.06 | - | 1.59 |
| Benzo(a)anthracene | 0.35 | 0.40 | 0.43 | 0.06 | 0.17 | 0.09 | 0.15 | 1.15 |
| Chrysene | 0.32 | 0.37 | 0.40 | - | 0.11 | - | 0.14 | 0.87 |
| Benzo(b or k)fluoranthene | - | - | - | - | 1.52 | 0.98 | 0.74 | 1.00 |
| Benzo(a)pyrene | 0.05 | 0.03 | 0.06 | 0.04 | 1.74 | 1.36 | 1.02 | 1.14 |
| Dibenzo(a,h)anthracene | 0.49 | 0.17 | 0.21 | 0.11 | 3.79 | 3.91 | 1.43 | 1.28 |
| Benzo(ghi)perylene | - | - | - | - | 2.48 | 2.96 | 1.12 | 0.29 |
| Indeno(1,2,3-cd)pyrene | 0.16 | 0.05 | 0.08 | 0.05 | 1.73 | 2.06 | 0.78 | 0.97 |
| Bis(2-chloroisopropyl)ether | - | - | - | 0.55 | - | 0.60 | - | - |
| N-Nitrosodi-n-propylamine | 0.13 | 0.15 | 0.06 | 0.08 | 0.06 | 0.06 | 0.03 | 0.23 |
| Bis(2-chloroetoxy)methane | - | - | - | - | - | 0.05 | 0.03 | 0.10 |
| N-Nitrosodiphenylamine | - | - | - | - | - | 14.71 | - | 0.55 |
| Diethylphthalate | 1.3 | 1.6 | 1.4 | 10.3 | 1.1 | 14.3 | 1.3 | 47.2 |
| Azobenzene | - | - | - | - | - | 0.43 | 0.05 | 0.07 |
| Carbazole | - | - | - | - | - | 4.01 | 0.06 | 0.26 |
| Di-n-butylphthalate | 2.18 | 2.52 | 3.53 | 3.11 | 0.90 | 4.08 | 0.54 | 15.16 |
| Butylbenzylphthalate | 0.19 | 0.15 | 0.22 | 4.17 | 0.98 | 4.29 | 1.25 | 8.84 |
| Bis(2-ethylhexyl)phthalate | 2.4 | 3.3 | 6.3 | 8.8 | 2.7 | 27.7 | 4.9 | 32.8 |
| Di-n-octylphthalate | 0.43 | 0.52 | 0.04 | 0.06 | 2.26 | 2.47 | 1.26 | 2.08 |
| 1,3-Dichlorobenzene | 0.17 | 0.19 | 0.23 | 0.50 | 0.13 | 0.17 | 0.10 | 0.16 |
| 1,4-Dichlorobenzene | 0.17 | 0.19 | 0.23 | 0.51 | 0.14 | 0.17 | 0.10 | 0.17 |
| 1,2-Dichlorobenzene | 0.16 | 0.17 | 0.21 | 0.46 | 0.12 | 0.16 | 0.09 | 0.16 |
| Hexachlorethane | - | - | - | - | - | 0.39 | - | 0.60 |
| Nitrobenzene | - | - | - | 6.87 | - | 0.33 | - | 0.24 |
| Isophorone | - | - | - | 6.87 | - | 0.33 | - | 0.24 |
| 1,2,3-Trichlorobenzene | - | - | - | 0.91 | - | - | 0.47 | 0.04 |
| 2,6-Dinitrotoluene | 0.60 | - | 1.05 | 0.55 | 5.57 | 0.60 | - | 1.05 |
| 2,4-Dinitrotoluene | 0.26 | - | 2.95 | 0.09 | 1.04 | 0.26 | - | 2.95 |
| alpha-BHC | - | - | 6.21 | - | - | - | - | 6.21 |
| gamma-BHC | - | - | 0.44 | - | 1.08 | 0.24 | - | 0.44 |
| beta-BHC | - | - | 0.59 | - | 0.83 | 1.31 | - | 0.59 |
| delta-BHC | - | - | 0.38 | - | 2.14 | 0.19 | - | 0.38 |
| Heptachlor | 0.57 | - | - | - | 7.69 | 0.57 | - | - |
| Endosulfan1 | - | - | - | - | - | 0.21 | 0.22 | 0.77 |
| 4,4-DDE | - | - | - | - | - | 0.04 | 0.03 | 0.02 |
| Endrin | - | - | - | - | 0.33 | 0.31 | 0.33 | 0.66 |
| Endosulfan2 | - | 0.56 | 0.13 | 0.57 | 0.14 | 0.15 | 0.52 | 1.16 |
| 4,4-DDT | - | - | - | - | - | - | 0.15 | 0.15 |
| Endosulfan sulfate | - | - | - | - | 0.15 | 0.12 | 0.26 | 0.70 |
| Methoxychlor | - | - | - | 0.34 | 0.28 | 1.31 | 0.20 | 2.78 |
| Endrin ketone | - | - | - | - | - | - | - | 0.67 |
| 2,2',3,3',4,4',6-PCB | - | - | - | - | - | - | - | 0.44 |
| Hydrocarbon C10 | 0.14 | 0.16 | 0.15 | 0.17 | 0.11 | 0.17 | <0.1 | 0.23 |
| Hydrocarbon C11 | 0.28 | 0.34 | 0.34 | 0.26 | 0.18 | 0.55 | 0.16 | 0.77 |
| Hydrocarbon C12 | 0.89 | 0.98 | 0.93 | 0.86 | 0.44 | 0.53 | 0.31 | 0.30 |
| Hydrocarbon C14 | 0.21 | 0.21 | 0.18 | 0.28 |  | 0.42 | 0.11 | 0.93 |
| Hydrocarbon C16 | 0.25 | 0.24 | 0.18 | 0.25 | <0.1 | 0.49 | <0.1 | 1.37 |
| Hydrocarbon C17 | 0.46 | <0.1 | <0.1 | 0.94 | 0.28 | 2.36 | 0.21 | 7.43 |
| Hydrocarbon C18 | 0.11 | <0.1 | 0.13 | 0.35 | <0.1 | 0.83 | <0.1 | 2.81 |
| Hydrocarbon C10 | 0.14 | 0.16 | 0.15 | 0.17 | 0.11 | 0.17 | <0.1 | 0.23 |
| Hydrocarbon C20 | - | - | - | 0.27 | - | 0.63 | - | 3.71 |
| Hydrocarbon C24 | - | - | - | - | 0.21 | 0.38 | - | 3.82 |
| Hydrocarbon C26 | 0.24 | - | 0.61 | 0.23 | 3.87 | 2.34 | 0.81 | 13.67 |
| Hydrocarbon C28 | 0.10 | 0.17 | 0.34 | 0.31 | 4.97 | 4.12 | 1.67 | 10.65 |
| Hydrocarbon C30 | - | - | 0.14 | 0.25 | 1.51 | 2.46 | 0.36 | 5.94 |
| Hydrocarbon C32 | - | - | - | 0.24 | 0.82 | 2.91 | 0.16 | 3.11 |
| Hydrocarbon C34 | - | - | - | 0.25 | 0.36 | 1.92 | - | 2.56 |

Annex 10: 	Sampling campaign to the area of Milano. 23.10.2001. Concentrations for the volatile contaminants measured by headspace-GC-MS, [μg/l].

| Compound | ## NM1 | NM2 | NM3 | ## NM4 | NM5 | NM6 | NM7 | NM8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Chloroform | 0.83 | - | - | - | - | - | - | 1.74 |
| Trichloroethylene | - | - | - | - | - | 1.85 | - | 3.53 |
| Tetrachloroethylene | 0.72 | - | - | 1.49 | - | - | 0.51 | - |
| 1,2,4-Trimethylbenzene | - | - | - | - | - | 0.66 | - | 1.65 |
| p-Isopropyltoluene | - | - | - | - | - | 0.65 | - | 1.34 |
| n-Butylbenzene | - | - | - | - | - | - | - | 0.52 |
| Naphthalene | - | - | - | - | - | - | - | 0.90 |

Annex 11: 	Sampling campaign to the River Seveso. 12.02.2002. Concentrations for the volatile contaminants measured by headspace-GC-MS, [μg/l].

| Compound | ## RS1 | RS2 | RS3 | ## RS4 | RS5 | RS6 | RS7 | RS8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Chloroform | - | - | - | 29.3 | 16.1 | - | - | - |
| Toluene | 0.11 | 0.26 | 0.19 | 0.12 | - | - | - | - |
| Tetrachloroethylene | 0.35 | 7.49 | 0.45 | 0.35 | 0.91 | - | 0.17 | 0.26 |
| Trichloroethylene | - | 0.12 | - | - | - | - | - | - |
| Ethylacetate | - | - | - | - | - | - | 0.91 | 0.37 |
| Trimethylsilanol | - | - | - | - | - | 0.24 | - | - |

# Acknowledgements

The C18-SPE GC-MS analyses were performed by Roberto Tilio.

The headspace GC-MS and the physical and sum parameters analyses were done by Elisabetta Canutti.

We want to thank Lariana Depur SpA (Fino Mornasco) for the kind cooperation.

# References

[1]	EU: Council Directive 96/61/EC of 24.09.1996 concerning integrated pollution 	prevention and control, Off. J. Europ. Commun. 1996, L 257, 10.10.1996, p. 26-40.

[2]	EU: Decision Establishing the List of Priority Substances in the Field of Water Policy 	and Amending Directive 2000/60/EC, Off. J. Europ. Commun. 2001, Decision No. 	2455/2001/EC (20. Nov. 2001), Annex X to Directive 2000/60/EC.

[3]	EU: Water Framework Directive, Off. J. Europ. Commun. 2000, Directive 	2000/60/EC, 23. Oct. 2000.

[4]	W. M. A. Niessen: State-of-the-Art in Liquid Chromatography-Mass Spectrometry, J. 	Chromatogr. A 1999, 856, 179-197.

[5]	S. D. Richardson: Environmental Mass Spectrometry, Anal. Chem. 2000, 72, 4477-	4496.

[6]	T. Reemtsma: The Use of Liquid Chromatography-Atmospheric Pressure Ionization-	Mass Spectrometry in Water Analysis - Part I: Achievements, Trends in Anal. Chem. 	2001, 20, 500-517.

[7]	S. D. Richardson: Mass Spectrometry in Environmental Sciences, Chem. Rev. 2001, 	101, 211-254.

[8]	D. W. Kolpin; E. T. Furlong; M. T. Meyer; E. M. Thurman; S. D. Zaugg; L. B. Barber; 	H. T. Buxton: Pharmaceuticals, Hormones, and Other Organic Wastewater 	Contaminants in U.S. Streams, 1999-2000: A National Reconnaissance, Environ. Sci. 	Technol. 2002, 36, 1202-1211.

[9]	T. Reemtsma: The Use of Liquid Chromatography-Atmospheric Pressure Ionization-	Mass Spectrometry in Water Analysis - Part II: Obstacles, Trends in Anal. Chem. 	2001, 20, 533-542.

[10]	M. C. Hennion: Solid-Phase Extraction: Method Development, Sorbents, and 	Coupling with Liquid Chromatography, J. Chromatogr. A 1999, 856, 3-54.

[11]	M.-C. Hennion; V. Pichon: Solid-Phase Extraction of Polar Organic Pollutants from 	Water, Environ. Sci. Technol. 1994, 28, 576A-583A.

[12]	P. L. Ferguson; C. R. Iden; B. J. Brownawell: Analysis of Alkylphenol Ethoxylate 	Metabolites in the Aquatic Environment Using Liquid Chromatography-Electrospray 	Mass Spectrometry, Anal. Chem. 2000, 72, 4322-4330.

[13]	C. Sonnenschein; A. M. Soto: An Updated Review of Environmental Estrogen and 	Androgen Mimics and Antagonists, J. Steroid Biochem. Molec. Biol. 1998, 65, 143-	150.

[14]	R. Trussel: Endocrine Disrupters and the Water Industry, J. Am. Wat. Wks Assoc. 	2001, Feb., 58-65.

[15]	T. Jones-Lepp; C. L. Gerlach; E. J. Cooter: The Power of Analytical Methods for 	Measuring Suspected Endocrine Disrupting Compounds: A Pilot Field Study, Trends 	in Anal. Chem. 2000, 19, 286-291.

[16]	T. Reemtsma: Analysis of Sulfophthalimide and Some of its Derivatives by Liquid 	Chromatography-Electrospray Ionization Tandem Mass Spectrometry, J. 	Chromatogr. A 2001, 919, 289-297.

[17]	A. Cohen; K. Klint; S. Bowadt; P. Persson; J. A. Joensson: Routine Analysis of 	Alcohol and Nonylphenol Polyethoxylates in Wastewater and Sludge Using Liquid 	Chromatography-Electrospray Mass Spectrometry, J. Chromatogr. A 2001, 927, 103-	110.

[18]	M. Takino; S. Daishima; K. Yamaguchi: Determination of Nonylphenol Ethoxylate 	Oligomers by Liquid Chromatography-Electrospray Mass Spectrometry in River 	Water and  Non-Ionic Surfactants, J. Chromatogr. A 2000, 904, 65-72.

[19]	DIN: Guidelines for the Determination of Total Organic Carbon (TOC) and Dissolved 	Organic Carbon (DOC), DIN EN 1484, 1997.

[20]	F. M. Lagerwerf; W. D. Van Dongen; R. J. J. M. Steenvoorden; M. Honing; J. H. G. 	Jonkman: Exploring the Boundaries of Bioanalytical Quantitative LC-MS/MS, Trends 	in Anal. Chem. 2000, 19, 418-427.

[21]	M. Petrovic; A. Diaz; F. Ventura; D. Barceló: Simultaneous Determination of 	Halogenated Derivatives of Alkylphenol Ethoxylates and their Metabolites in Sludges, 	River Sediments and Surface, Drinking and Wastewater by LC-MS, Anal. Chem. 	2001, 73, 5886-5895.

[22]	P. L. Ferguson; C. R. Iden; B. J. Brownawell: Analysis of Nonylphenol and 	Nonylphenol Ethoxylates in Environmental Samples by Mixed-Mode High-	Performance Liquid Chromatography-Electrospray Mass Spectrometry, J. 	Chromatogr. A 2001, 938, 79-91.

[23]	R. J. M. Vervoort; E. Ruyter; A. J. J. Debets; H. A. Claessens; C. A. Cramers; G. J. 	De Jong: Characterisation of Reversed-Phase Liquid Chromatography Stationary 	Phases for the Analysis of Basic Pharmaceuticals: Eluent Properties and Comparison 	of Empirical Test Methods, J. Chromatogr. A 2001, 931, 67-79.

[24]	H. Boernick; T. Grischek; E. Worch: Determination of Aromatic Amines in Surface 	Waters and Comparison of Their Behaviour in HPLC and on Sediment Columns, 	Fresenius J. Anal. Chem. 2001, 371, 607-613.

[25]	O. Jauregui; E. Moyano; M. T. Galceran: Liquid Chromatography-Atmospheric 	Pressure Chemical Ionization Mass Spectrometry for Chlorinated Phenolic 	Compounds. Application to the Analysis of Polluted Soils, J. Chromatogr. A 1998, 	823, 241-248.

[26]	D. Puig; D. Barceló: Comparison of Different Sorbent Materials for On-Line Liquid-	Solid Extraction Followed by Liquid Chromatographic Determination of Priority 	Phenolic Compounds in Environmental Waters, J. Chromatogr. A 1996, 733, 371-	381.

[27]	R. Wissiack; E. Rosenberg; M. Grasserbauer: Comparison of Different Sorbent 	Materials for On-Line Solid-Phase Extraction With Liquid-Chromatography-	Atmospheric Pressure Chemical Ionization Mass Spectrometry of Phenols, J. 	Chromatogr. A 2000, 896, 159-170.

[28]	S. Geiss; J. W. Einax: Comparison of Detection Limits in Environmental Analysis - Is 	it Possible? An Approach on Quality Assurance in the Lower Working Range by 	Verification, Fresenius J. Anal. Chem. 2001, 370, 673-678.

[29]	R. Loos; G. Hanke: Determination of Endocrine Disrupting Alkylphenols in Water by 	Solid-Phase Extraction and Liquid Chromatography - Electrospray Ionization - Mass 	Spectrometry Detection, J. Chromatogr. A 2002, submitted.

[30]	B. K. Matuszewski; M. L. Constanzer; C. M. Chavez-Eng: Matrix Effect in 	Quantitative LC-MS/MS Analyses of Biological Fluids, Anal. Chem. 1998, 70, 882-	889.

[31]	R. King; R. Bonfiglio; C. Fernandez-Metzler; C. Miller-Stein; T. Olah: Mechanistic 	Investigation of Ionization Suppression in Electrospray Ionization, J. Am. Soc. Mass 	Spectrom. 2000, 11, 942-950.

[32]	M. Petrovic; D. Barceló: Analysis of Ethoxylated Nonionic Surfactants and their 	Metabolites by LC-API-MS, J. Mass Spectrom. 2001, 36, 1173-1185.

CAS: Chemical Abstract Services

EU number: European Inventory of Existing Commercial Chemical Substances (EINECS) or European List of Notified Chemical Substances (ELINCS)

n.a.	not applicable